❶ Reference Standard：ICHQ1、Chinese Pharmacopoeia 2025 edition; 
❷ Insulation material: Overall high-density polyurethane foam technology, with good insulation and moisture retention performance; 
❸ Chamber Materials:The exterior material is coated with high quality steel plate,The inner material is made of stainless steel 304, no pollution source, easy to clean; 
❹ Control system: Programmable color touch screen controller, more than 7 inches, display frequency conversion output data; 
❺ Three Level Authority：Adopt the user name and password to log in. Different user names  have different permissions at least three levels or above; 
❻ Audit tracking: The control system has historical alarm recording and operation recording functions, and supports exporting in an unchangeable file format using a USB flash drive; 
❼ Refrigeration system: Imported fully enclosed compressor with frequency conversion, save more than 50% of electricity and water; 
❽ Humidity control: Original imported capacitive humidity sensor, high precision, low drift, long service life, maintenance free; 
❾ Data management: configuring needle micro printers and electronic data storage functions; 
❿ Safety device: Compressor overheat and over pressure overload protection, water shortage protection.

Anti dry burning protection system, independent over temperature protection alarm system; 
⓫ Alarm system：On-site audible & visual alarm; 
⓬ Double Door structure：Interior door tempered glass door，Open the outer door to observe the sample.

The temperature and humidity inside the container will not change in a short time. The outer door is solid can keep temperature and humidity.

It can also prevent the influence of external light; 
⓭ Controlling Temperature Precision：Temperature Fluctuation＜±0.5℃；Temperature Uniformity ＜±1.0℃ ; 
⓮ Controlling Humidity Precision：Humidity Fluctuation ＜ ±2%RH；Humidity Uniformity ＜  ±3%RH; 
⓯ Power：AC220V±10% 50Hz/60Hz; 
⓰ Environment Temperature：+5  ~  35℃ ; 
⓱ Optional: Temperature- humidity deviation, sudden power failure, remote SMS alarm.

| Model | Temperature Range( ℃ ) | Humidity Range (RH) | Capacity (L) | Interior Dimensions W×D×H(mm) | External Dimensions W×D×H(mm) | Power (kW) | Shelf (Standard) | Remarks |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Labonce-150GS-FC | 10~65 | 15~95% | 150 | 600×405×620 | 780×830×1480 | 15 | 3/6 | Energy saving > 50%, Water saving > 80% |
| Labonce-250GS-FC | 10~65 | 15~95% | 250 | 600×500×830 | 780×880×1650 | 1.8 | 3/7 |  |
| Labonce-400GS-FC | 10~65 | 15~95% | 400 | 600×700×950 | 780×1060×1750 | 2.0 | 3/10 |  |
| Labonce-500GS-FC | 10~65 | 15~95% | 500 | 680×680×1080 | 860×1050×1850 | 2.2 | 4/11 |  |